Synchrotron Radiation Monitors for LHC

R. Jung, L. Ponce

Ackowledgements to: 	J. Bosser, F. Méot

P. Komorowski, D. Tommasini, R. van Weelderen

R. Ostojic, S. Russenschuck, T. Taylor

F. Bordry, P. Burla, M. Sassowsky, L. Tavian

Synchrotron Radiation (SR) profile monitors for LHC

# At the time of the Emittance workshop in July 2000:

the main profile monitor was the SR monitor of the D2 bending magnet in IR5 able to work from 2 TeV to 7 TeV, best at collision optics, supplemented from 450 GeV to 2 TeV by three candidates: a Wiggler SR monitor, an Ionisation Profile Monitor and a Luminescence monitor, all in IR4.

In Summer 2001, the Layout version 6.4 introduced a change of the IR4 layout where we can now combine an Undulator with the D3 bending magnet which makes it possible to use this monitor from 450 GeV to 7 TeV by taking first the light generated by the Undulator and later in the ramp the light generated in the same direction by the edge of D3.

We consider now this monitor as the main profile monitor, supplemented by:

- the IR5 D2 SR monitor, if the performance in collision optics at 7 TeV has to be maximised,

- the gas monitors in IP4 for redundancy in proton mode and for ions.

“SC Undulator In D3” profile monitor in IR4

Undulator: main characteristics

An Undulator:

- is a periodic magnetic structure  - + - +  with N periods

- has a principal direction along its axis

- concentrates the SR in a forward cone around this direction through interference

- emits SR with an angular spectral energy distribution proportional to     B2 N [sin k N/kN]2

- emits wavelengths which are a function of beam energy, undulator period and angle wrt ppal direction

#### IR4 Undulator SR monitor performance (1)

We have chosen an undulator with 2 periods of  28 cm to generate a maximum of SR at 450 GeV and enough on the ramp until D3 takes over around 2 TeV.

450 GeV

28cm period Undulator SR characteristics

1 TeV

#### IR4  D3/Undulator SR Monitor performance (2)

#### The monitor taking the light from the Undulator and from the D3 edge is sensitive to the positions of the two light sources.

#### Two effects can be seen:

#### an effect on the size of the image of the beam core

#### optics generated tails

#### To minimise both effects, the source separation has to be kept below ~1 m.

Undulator in D3: main issues

The undulator will have a field of ~ 6 T and will introduce a 0.1mm orbit wiggle in the worst case.

It has to be powered with 200A.

#### The D3s are manufactured by BNL.

We propose to locate the Undulator as close as possible to D3 to minimise the negative side effects.

Due to this, the Undulator has to be incorporated into the cryostat of D3, which will be an activity under the responsibility of CERN.

There are in fact two solutions:

- The minimal one is to integrate the Undulator into the D3 cryostat, before the D3 cold mass. In this case, a study has shown that the perturbations to the cold mass of D3 are negligible, which minimises the intervention of BNL, while the separation can be kept around 80cm

- The other solution would be to integrate the Undulator into the cold mass of D3, which would further decrease the separation but needs more effort from BNL.

Both solutions are being discussed by the LHC project management with BNL.

#### Vacuum Chamber

#### At the exit of D3 and up to the extraction mirror, the vacuum chamber will be an off-axis cylinder with a diameter increased to 90mm to transmit the SR. There will be tapered transitions at both ends to minimise the vacuum chamber impedance due to this change in cross-section.

The light will be intercepted horizontally by a movable mirror and brought out vertically through a vacuum window to the SR telescope.

#### SR Telescope

#### The optical characteristics of the monitor are similar to those of the LEP monitors, so it is proposed to adapt the LEP telescope BEUV to the LHC requirements. (There are still 2 LEP telescopes available!).

#### The telescope is built on a 3.2m optical bench housed inside a 4m stainless steel tube.

The main changes will consist in replacing the small Be extraction mirror by a 50 x 50 mm2 mirror and the deformable first folding mirror by a normal mirror.

#### The movable camera can be replaced by available CCD/Intensifier or rotating Multi-anode PM set-ups to do single turn/single bunch and time integrated imaging.

#### If an additional magnification is asked for at 7 TeV, it can be made available.

Tails can be measured with the “corona” technique developped in LEP.

#### Detectors and Signal Processing

#### In IP4 where the radiation level should be small (<10 Gy/Y), CCDs can be used.

#### MCP intensifiers can provide:

#### the high sensitivity needed for single turn measurements down to pilot pulse intensity

#### the pulsed operation for bunch-by-bunch and turn-by-turn measurements.

#### Technology available today in BI can provide:

#### light intensity gains up to 104

#### 2-Dimensional profile measurements up to 14 turns of a given bunch at a repetition rate up to 11kHz

- projections at 40MHz rate with a 32-anode PM

- tail measurements down to 10-3 at least on time integrated (20ms or more) signals

#### Cross calibration and Alignment set-up

#### The reference device for profile measurements is the Wire Scanner.

The Wire Scanner is limited in intensity, but as the SR monitor is not intensity dependent in the LHC SR power range, there is no problem to check the monitor calibration at low intensity.

#### From LEP we have learned that the Wire Scanner has to be closest to the device to calibrate to avoid any error from transporting the beam size from one monitor to the other because of  beating.

#### For that reason the wire scanner will be located at the exit of D3.

#### For precision alignment and future maintenance and modifications, it is important to provide a specific alignment tool, in order to provide an alignment precision better than 0.1 mrad.

#### This will be done with no beam in LHC with a laser beam aimed at a distance of ~10m at the input hole (±1mm) of a precision folding mirror set-up which is located on the SR beam path and sends the light to the extraction mirror. These tanks are available from LEP where they have proven their usefulness.

#### IP5 D2 bending magnet SR Monitor

#### The IP5 monitor will be of similar design than the IP4 monitor.

#### In IP5, the light will be generated in the D2 magnet, and as this monitor is expected to be used only at high energy (above 2 to 4 TeV), it will not need the radiation emitted at the edge, but can make use of the light generated inside the D2 magnet, which makes it possible to prevent it from parasitic light coming from D1 by a slit at the focal point of the focusing mirror.

Because of the change in  when the beams are colliding, the beam sizes will then be 2 to 2.5 times larger than in the IR4 monitors, which can be interesting if the best resolution is looked for.

Tail measurements will also be more precise at high energy.

The disadvantage of this monitor location is the higher radiation levels of 10 to 100 gy/Y or more which will put restrictions on the use of CCD detectors and on the maintenance operations.

In the cost reduction exercise of BI, we have proposed to delay the construction of this monitor to some later moment.

## Spatial Sampling / resolution

The beam size has a large excursion at the IR4 monitor, from 13 pixels*/ at injection to 3.3 px/ at 7 TeV.

The magnification has been optimised there for making single turn measurements of the pilot bunch at injection with a precision of ~ 6%.

#### If more pixels per  are requested at top energy, an additional magnification can be introduced or the sensitivity/precision can be degraded for the pilot pulse at 450 GeV.

#### With 9 px/ at 7 TeV in collision optics and enough signal to saturate the central pixel of a CCD, it will be possible to make high precision measurements with the D2 monitor for colliding beams.

* pixel size brought back to the source: ~ 100m

#### Monitor sensitivities : Single turn

There is a difference of a factor 2 to 3 on beam size at 7 TeV in collision optics between the Undulator and the D2 monitors.

*   CCDs saturate at ~ 2 x 106 photons/px

1000 primaries are requested (statistics) on the central pixel to measure a profile to 3%

200 primaries are needed to use the full range of the final detector after a light amplification of 104

#### +   at injection optics

++ at collision optics

#### Monitor sensitivities: TV mode  (225 revolutions)

## There is enough signal to

## do precision measurements [% range] with single bunches

## measure tails of single bunches at high energy and of batches at low energy with the “corona” method.

* CCDs saturate at ~ 2 x 106 photons/px

1000 primaries are requested on the central pixel to measure a profile to 3%

+  at injection optics

++ at collision optics

Precision limitation

The main limitations for the measurement of the profiles will be:

Diffraction: 	due to the lower beam , this will be less severe than in LEP

Depth of field:	this effect will be more important in D2 than in LEP because of the larger bending radius

Parasitic light:	~ less a problem than in LEP because of the low light levels (detection level cut-off)

The diffraction and depth of field effects can be calculated and later measured, and the results checked in cross-calibration runs with the Wire Scanners as has been demonstrated in LEP:

Pb 82+   IONS

Preliminary thoughts have been given to assess the possibility to monitor ion beams with SR.

Because of their lower , the fundamental Undulator wavelength is shifted into the Infra Red at injection, and hence the usual detectors cannot be used.

To overcome this, two possibilities have been considered:

- use Infra-red detectors which can be found in the range 1000-3500nm where the vacuum window is transparent

- add a period (or two) of 10cm in front of the proton Undulator to shift the wavelengths down towards the visible. A field of 2.4T seems feasible.

With both options, the present calculations show that it should be possible to observe the beams and measure profiles in TV mode at injection.

At higher beam energies, there seems to be no problem.

Tests will be made in 2002 during the SPS  Pb 82+  run to check the theoretical calculations.

E2000  SL AP adjusted demands on profile precision

- Injection matching: 	10% t-b-t to limit the filamentation blow-up to ~ 1%

non-interceptive monitoring during filling as “watchdog”

- Circulating beam: real time TV, gating for single bunch but not turn-by-turn

4% relative, or 4% absolute if cross-calibrated with Wire Scanner

5-10% precision for a scan of the injected batch before next injection

5-10% precision for coherent modes, b-b-b, and t-b-t if possible

- Tails: sensitive to 10-4  of bunch with ~10% accuracy on measurement

no time requirement

- Ions: 10% precision for 20ms integration

- Pilot: TV [20ms] observation

All demands can be ~ satisfied with the proposed set-up and technology available in BI.

Further work to be done

Finalise the Undulator design and its integration in the D3 cryostat together with BNL.

Study the case of Pb82+ completely, perform tests in the SPS.

Determine precisely the influence of diffraction and depth of field on beam size broadening.

Finalise the optics design.

Follow / participate in new detector developments:	radiation hardness

faster readout

higher dynamic range.

Conclusion

All “reasonable” demands can be satisfied for protons with our available technology.

More work has to be done for ions at injection energy.

Some more discussions/investigations have to be done on the subject of 4 to 5% precisions.

# Text boxes

|  | Beam Energy | 450 GeV single turn | 7 TeV single turn ++ |
| --- | --- | --- | --- |
|  | Beam Intensity | 5.E+09 | 1.1E+11 | 5.E+09 | 1.1E+11 |
| Max “Signal” * / "pixel" of | /3 at | for primary |  |  |  |  |
| D3 & SC Undulator [IR4] | 7 TeV | photons | 3.3E+02 | 7.3E+03 | 6.2E+05 | 1.4E+07 |
| BSRT [D2 in IR5]: 2 - 7 TeV | 7 TeV + | photons |  |  | 2.8E+04 | 6.1E+05 |
| OTR | 450 GeV | photons | 8.0E+04 | 1.6E+06 |  |  |

|  | Beam Energy | 450 GeV 20ms [TV] | 7 TeV 20ms [TV] ++ |
| --- | --- | --- | --- |
|  | Beam Intensity | 5.E+09 | 1.1E+11 | 5.E+09 | 1.1E+11 |
| Max “Signal” * / "pixel" of | /3 at | for primary |  |  |  |  |
| D3 & SC Undulator [IR4] | 7 TeV | photons | 7.5E+04 | 1.6E+06 | 1.4E+08 | 3.1E+09 |
| BSRT [D2 in IR5] | 7 TeV + | photons |  |  | 2.4E+07 | 5.3E+08 |
| OTR | 450 GeV | photons |  |  |  |  |